## Schedule of Requirements
STRANFIELD TRADING COMPANY ¦ SCHEDULE OF REQUIREMENTS & PROCUREMENT SPECIFICATIONS
Document Ref: STC-PROC-2025-01 ¦ Target Destination: Lagos, Nigeria (Direct Air Dispatch) ¦ Issue Date: October 2025
Instructions to Vendors: Review technical specifications and required compliance standards below. Complete pricing placeholders in USD/GBP/EUR (indicate Incoterms 2020) and submit accompanying PQQ matrices (PQQ1 to PQQ5).
Item No. | Category & Description | Technical Specs & Compliance | MOQ | Unit Wholesale Price (Ex-Works / FCA / CIP)
1 | Medical-Grade Non-Contact Infrared Handheld Thermometers (Diagnostic & Screening Equipment) | • Clinical Accuracy: ±0.2°C (within 35.0°C to 42.0°C range) • Measuring Time: < 1 second; Optical distance: 3 - 5 cm • Display: High-definition backlit LCD with fever alarm • Regulatory: CE Medical Class IIa, ISO 13485:2016, FDA 510(k) Cleared | 500 Units | [Vendor to Quote - Specify Currency & Incoterm]
2 | Walk-Through Mass Thermal Imaging Camera System with Blackbody (Mass Screening & Surveillance) | • Uncooled focal plane thermal detector; Real-time AI multi-target face tracking • Calibration: High-precision external blackbody source (±0.3°C accuracy) • Standard: IEC 80601-2-59 compliant for human febrile screening • Inclusions: Optical/thermal dual-lens camera, blackbody, dual tripods & management software | 5 Systems | [Vendor to Quote - Specify Currency & Incoterm]
3 | Walk-Through Ultraviolet (UV) Disinfection Booth (Sanitization & Infection Control) | • Chamber: Automated multi-emitter UV-C portal (wavelength ≥254 nm) • Operation: 360° rapid surface decontamination cycle with motion activation • Safety: Optical shielding interlocks & instant emergency shutoff • Compliance: CE Medical / CE Machinery Directive & ISO 13485 | 2 Systems | [Vendor to Quote - Specify Currency & Incoterm]
4 | WHO VHF Category III Type 3-B / 4-B Biohazard Coveralls (Personal Protective Equipment - PPE) | • Protection: Category III Type 3-B (liquid-tight) & Type 4-B (spray-tight) • Biohazard Barrier: EN 14126:2003 certified (ISO 16604 Class 6 viral penetration resistance) • Construction: 100% impervious multi-layer fabric, sealed/taped seams, integrated storm flap • Design: Elasticated hood, wrists, waist, and ankles; thumb loops | 2,500 Units | [Vendor to Quote - Specify Currency & Incoterm]
5 | Heavy-Duty Biohazard Mortuary Body Bags (250+ Micron) (Mortuary & Biosecurity Supplies) | • Material: Heavy-duty virgin PE/PVC fluid-impermeable matrix (≥250 µm thickness) • Protocol: Strictly compliant with WHO Viral Haemorrhagic Fever (VHF) safe-handling protocols • Structural Capacity: Certified minimum 150 kg payload capacity • Features: 6 reinforced heavy-duty carry handles, curved U-shape dual zipper with fluid barrier flap | 1,000 Units | [Vendor to Quote - Specify Currency & Incoterm]
6 | Sodium Dichloroisocyanurate (NaDCC) Chlorine Tablets (5g) (Disinfection & Chemical Agents) | • Formulation: Rapidly effervescent 5g NaDCC disinfectant tablets • Concentration: Formulated for WHO 0.5% (5,000 ppm) available chlorine biocidal solution • Efficacy: EN 14476 virucidal efficacy against enveloped/non-enveloped viruses; EN 13727 bactericidal • Packaging & Stability: Airtight moisture-proof container; minimum 3-year shelf life | 50,000 Tablets (500 Tubs of 100) | [Vendor to Quote - Specify Currency & Incoterm]
7 | FFP3 / N95 Medical-Grade Biohazard Respirator Masks (Respiratory Protective Equipment) | • Filtration: ≥99% airborne particulate filtration down to 0.3 µm (FFP3 / N99 performance) • Standards: EN 149:2001+A1:2009 FFP3 NR, CE Certified (Regulation EU 2016/425) / NIOSH N95 • Resistance: Fluid-resistant outer layer conforming to WHO infection prevention standards • Fit: Adjustable padded nose bridge, ultrasonically welded dual head straps for secure seal | 10,000 Units | [Vendor to Quote - Specify Currency & Incoterm]

## PQQ1 - Quality & Regulatory
STRANFIELD TRADING COMPANY ¦ PQQ1: QUALITY ASSURANCE & REGULATORY COMPLIANCE MATRIX
Objective: Verify vendor regulatory credentials, manufacturing certifications, and clinical compliance standards. Vendors must state compliance status and provide certificate numbers, accredited issuing body details, and expiration dates.
Item No. & Equipment Name | Mandatory Certifications / Standards Required | Vendor Response & Certificate Reference Details
[Exemplar / Guidance] | [Mandatory regulatory benchmarks: CE Mark, ISO 13485:2016, ISO 9001, FDA 510(k), EN/IEC specific standards, WHO Pre-qualification / Guidelines] | [Vendor Guidance: State 'Fully Compliant' / 'Equivalent'. Enter: 1) Notified Body No. ¦ 2) Certificate / License ID ¦ 3) Expiry Date (DD/MM/YYYY) ¦ 4) Test Report Ref. Attach official PDF copies.]
Item 1: Medical-Grade Non-Contact Infrared Handheld Thermometers | • CE Medical Class IIa (Regulation EU 2017/745 MDR or MDD 93/42/EEC) • ISO 13485:2016 (Medical Devices Quality Management System) • FDA 510(k) Cleared or equivalent national medical device license • ASTM E1965-98 / IEC 60601-1-2 (Clinical accuracy & electrical safety) | [Enter: Compliance Status ¦ Certificate No. ¦ Notified Body Name & ID ¦ Expiry Date ¦ Attach Test Report]
Item 2: Walk-Through Mass Thermal Imaging Camera System with Blackbody | • IEC 80601-2-59:2017 (Screening thermographs for human febrile temperature) • ISO 9001 / ISO 13485 manufacturing quality certification • CE-EMC (EN 61000-6-3 / EN 61000-6-1) & CE-LVD (EN 62368-1) • External Blackbody calibration report traceable to national metrology institutes | [Enter: Compliance Status ¦ IEC 80601-2-59 Report No. ¦ Metrology Lab ID ¦ CE Certificate No. ¦ Expiry Date]
Item 3: Walk-Through Ultraviolet (UV) Disinfection Booth | • CE Medical / CE Machinery Directive (2006/42/EC) & Low Voltage Directive (2014/35/EU) • IEC 62471:2006 (Photobiological safety of lamps and lamp systems) • ISO 13485 or ISO 9001 QMS certification • Third-party laboratory microbiological efficacy report (log-4 / 99.99% kill rate) | [Enter: Compliance Status ¦ Safety Test Report No. ¦ UV-C Efficacy Lab Report Ref ¦ CE Certificate ID]
Item 4: WHO VHF Category III Type 3-B / 4-B Biohazard Coveralls | • Regulation (EU) 2016/425 Category III PPE (Module B + C2/D) • EN 14126:2003 + AC:2004 (Infective agents: ISO 16604 Class 6 synthetic blood/viral penetration) • EN 14605:2005+A1:2009 Type 3-B (Liquid-tight) & Type 4-B (Spray-tight) • EN ISO 13688:2013 (General requirements for protective clothing) | [Enter: Compliance Status ¦ EU Type-Examination No. ¦ 4-Digit Notified Body ID ¦ ISO 16604 Class Achieved]
Item 5: Heavy-Duty Biohazard Mortuary Body Bags (250+ Micron) | • WHO Infection Prevention and Control (IPC) protocol for Viral Haemorrhagic Fevers (VHF) • Material thickness certification: Virgin PE/PVC ≥250 µm (0.25mm) continuous gauge • Mechanical strength: ASTM D882 tensile / ISO 13934 tear resistance; 150 kg payload rating • Fluid containment: Seam burst pressure test & leak-proof certificate | [Enter: Compliance Status ¦ Material Gauge Lab Ref ¦ Load Test Certificate No. ¦ WHO VHF IPC Statement]
Item 6: Sodium Dichloroisocyanurate (NaDCC) Chlorine Tablets (5g) | • WHO Guideline on VHF Environmental Disinfection (delivers 0.5% / 5,000 ppm free chlorine) • EN 14476 (Virucidal quantitative suspension test for medical areas) • EN 13727 (Bactericidal) & EN 13624 (Fungicidal / Yeasticidal efficacy) • GMP (Good Manufacturing Practice) / ISO 9001 certified facility; EPA / ECHA registered | [Enter: Compliance Status ¦ EN 14476 Virucidal Report Ref ¦ GMP License No. ¦ EPA/ECHA Reg No. ¦ CoA Batch Ref]
Item 7: FFP3 / N95 Medical-Grade Biohazard Respirator Masks | • EN 149:2001+A1:2009 FFP3 NR under Regulation (EU) 2016/425 (Category III PPE) • NIOSH 42 CFR 84 N95 / FDA 510(k) cleared surgical respirator equivalent • Penetration of filter material: ≤1% leakage (≥99% filtration efficiency at 95 L/min) • Fluid resistance: EN 14683:2019 Type IIR (splash resistance ≥120 mmHg / 16.0 kPa) | [Enter: Compliance Status ¦ Notified Body / NIOSH TC-No. ¦ Filter Efficacy Test Ref ¦ CE Certificate Expiry]

## PQQ2 - Export & Direct Logistic
STRANFIELD TRADING COMPANY ¦ PQQ2: EXPORT & DIRECT LOGISTICS CAPABILITY MATRIX
Objective: Confirm vendor capability for direct international air freight dispatch to Murtala Muhammed International Airport (LOS), Lagos, Nigeria. Detail origin airport hubs, logistics partners, and lead times from purchase order to air dispatch.
Item No. & Equipment Name | Confirm Air Dispatch Capability (UK/EU/ZA to Lagos, NG) | Standard Dispatch Lead Time (Number of Days)
[Exemplar / Guidance] | [Vendor Guidance: Confirm YES/NO. State export origin airport hub (e.g. London Heathrow LHR, Frankfurt FRA, Amsterdam AMS, or Johannesburg JNB). State nominated air carrier/freight forwarder.] | [Vendor Guidance: Enter specific lead time in calendar days: 1) Production/Order Processing to Ex-Works handover; 2) Ex-Works to Airport Departure (AWB issue); 3) Flight Transit days to Lagos (LOS).]
Item 1: Medical-Grade Non-Contact Infrared Handheld Thermometers | [Confirm: YES/NO ¦ Origin Air Hub: LHR / FRA / AMS / JNB / Other ¦ Nominated Freight Forwarder ¦ Export Customs Clearance Capability: Confirmed] | [Specify Days: Ex-Stock dispatch: __ days ¦ Bulk Order (>2,000 units): __ days to air waybill issue ¦ Transit to LOS: __ days]
Item 2: Walk-Through Mass Thermal Imaging Camera System with Blackbody | [Confirm: YES/NO ¦ Origin Air Hub: LHR / FRA / AMS / JNB / Other ¦ Dedicated air cargo handling for sensitive optical equipment ¦ Transit Insurance: Included] | [Specify Days: Calibration & Pre-shipment QC: __ days ¦ Packing & Air Dispatch: __ days to AWB ¦ Transit to LOS: __ days]
Item 3: Walk-Through Ultraviolet (UV) Disinfection Booth | [Confirm: YES/NO ¦ Origin Air Hub: LHR / FRA / AMS / JNB / Other ¦ Oversized cargo pallet capability ¦ Dangerous goods / quartz tube packing compliant] | [Specify Days: Assembly/Testing: __ days ¦ Export Crate & Cargo Dispatch: __ days ¦ Transit to LOS: __ days]
Item 4: WHO VHF Category III Type 3-B / 4-B Biohazard Coveralls | [Confirm: YES/NO ¦ Origin Air Hub: LHR / FRA / AMS / JNB / Other ¦ Palletized air cargo capability ¦ Charter / Scheduled flight availability] | [Specify Days: Immediate stock dispatch: __ days ¦ Production lead time (if applicable): __ days to AWB ¦ Transit to LOS: __ days]
Item 5: Heavy-Duty Biohazard Mortuary Body Bags (250+ Micron) | [Confirm: YES/NO ¦ Origin Air Hub: LHR / FRA / AMS / JNB / Other ¦ Consolidated air freight or priority air express capability] | [Specify Days: Stock readiness: __ days ¦ Export dispatch to air carrier: __ days ¦ Transit to LOS: __ days]
Item 6: Sodium Dichloroisocyanurate (NaDCC) Chlorine Tablets (5g) | [Confirm: YES/NO ¦ Origin Air Hub: LHR / FRA / AMS / JNB / Other ¦ IATA Dangerous Goods (DG) certified air carrier & forwarder handling] | [Specify Days: DG Export declaration & packaging: __ days ¦ Handover to airline: __ days ¦ Transit to LOS: __ days]
Item 7: FFP3 / N95 Medical-Grade Biohazard Respirator Masks | [Confirm: YES/NO ¦ Origin Air Hub: LHR / FRA / AMS / JNB / Other ¦ High-volume volumetric air cargo booking capability] | [Specify Days: Warehouse dispatch: __ days ¦ Airport delivery & customs clearance: __ days ¦ Transit to LOS: __ days]

## PQQ3 - Packaging & Dimensions
STRANFIELD TRADING COMPANY ¦ PQQ3: PACKAGING SPECIFICATIONS & FREIGHT DIMENSIONS MATRIX
Objective: Capture detailed volumetric dimensions, carton packaging configurations, and gross weights to enable accurate air freight payload and cargo container planning for shipment to Lagos, Nigeria.
Item No. & Equipment Name | Master Carton / Unit Dimensions (L x W x H) | Gross Weight & Packaging Specifications
[Exemplar / Guidance] | [Vendor Guidance: Detail dimensions in cm or mm (Length x Width x Height). State packaging hierarchy: Primary unit box, Inner carton pack, Master export carton (e.g. 50 units/carton, 60 x 40 x 35 cm, 0.084 CBM).] | [Vendor Guidance: State Gross Weight (kg) and Net Weight (kg) per carton. Specify palletization: Cartons per pallet (Ti/Hi) and pallet total weight/height. Detail hazmat / IATA lithium battery labels if applicable.]
Item 1: Medical-Grade Non-Contact Infrared Handheld Thermometers | [Enter: Unit Box Dims (e.g. 18 x 11 x 5 cm) ¦ Master Carton Dims: __ cm x __ cm x __ cm ¦ Pack: 40 or 50 pcs/carton ¦ Volume per Carton: __ CBM] | [Enter: Carton Gross Weight: __ kg ¦ Net Weight: __ kg ¦ Battery: Alkaline (excluded) or Lithium (specify IATA DGR Section II / UN3481) ¦ Pallet Ti/Hi: __]
Item 2: Walk-Through Mass Thermal Imaging Camera System with Blackbody | [Enter: Crated Dimensions: Camera Module Box (__ x __ x __ cm), Blackbody Box (__ x __ x __ cm), Tripod Box (__ x __ x __ cm) ¦ Total Volume: __ CBM] | [Enter: Total System Gross Weight: __ kg ¦ Flight Case Specification: Heavy-duty aluminum/resin case with custom foam cutouts & shock indicators]
Item 3: Walk-Through Ultraviolet (UV) Disinfection Booth | [Enter: Shipping Crate Dimensions: __ cm (L) x __ cm (W) x __ cm (H) ¦ State whether shipped Knock-Down (Flat-Pack) or Assembled Portal ¦ Total Volume: __ CBM] | [Enter: Gross Crated Weight: __ kg ¦ Net Weight: __ kg ¦ Crate Spec: ISPM-15 heat-treated export timber with forklift skids & moisture barrier]
Item 4: WHO VHF Category III Type 3-B / 4-B Biohazard Coveralls | [Enter: Master Carton Dims: __ cm x __ cm x __ cm ¦ Pack: 50 pcs/carton (each individually vacuum-sealed in sterile polybag) ¦ Volume per Carton: __ CBM] | [Enter: Gross Weight: __ kg/carton ¦ Net Weight: __ kg/carton ¦ Pallet Configuration: __ cartons per Euro/Standard pallet ¦ 5-ply export corrugated]
Item 5: Heavy-Duty Biohazard Mortuary Body Bags (250+ Micron) | [Enter: Master Carton Dims: __ cm x __ cm x __ cm ¦ Pack: 10 or 20 bags/carton (individually wrapped with biohazard markings) ¦ Volume per Carton: __ CBM] | [Enter: Gross Weight: __ kg/carton (approx 16 - 22 kg) ¦ Net Weight: __ kg ¦ Heavy-duty weather-sealed export carton; strapping reinforced]
Item 6: Sodium Dichloroisocyanurate (NaDCC) Chlorine Tablets (5g) | [Enter: Master Carton Dims: __ cm x __ cm x __ cm ¦ Pack: e.g. 6 x 100-tablet tubs per carton (or 1,000-tab drums) ¦ Volume per Carton: __ CBM] | [Enter: Gross Weight: __ kg/carton ¦ Net Weight: __ kg ¦ IATA Dangerous Goods packaging: UN 2465 / UN 3077 certified, tamper-evident sealed HDPE tubs]
Item 7: FFP3 / N95 Medical-Grade Biohazard Respirator Masks | [Enter: Unit Box Dims: 20 pcs/box ¦ Master Carton Dims: __ cm x __ cm x __ cm (e.g. 20 boxes = 400 pcs/carton) ¦ Volume per Carton: __ CBM] | [Enter: Gross Weight: __ kg/carton ¦ Net Weight: __ kg ¦ Export 5-ply double-wall corrugated carton, shrink-wrapped palletization]

## PQQ4 - Financials & MOQs
STRANFIELD TRADING COMPANY ¦ PQQ4: COMMERCIAL TERMS, MOQS & TIERED PRICING MATRIX
Objective: Establish vendor minimum bulk order quantities (MOQs), tiered volume pricing discounts, agreed Incoterms 2020 delivery points, and standard international payment terms for emergency and planned procurement.
Item No. & Equipment Name | Minimum Bulk Order Quantity (MOQ) | Volume Price Breaks / Tier Pricing | Incoterms 2020 & Payment Terms Offered
[Exemplar / Guidance] | [Vendor Guidance: State baseline MOQ in units/packs. Indicate whether smaller trial/air-express emergency batches are accepted.] | [Vendor Guidance: Detail unit prices per tier in USD/GBP/EUR: Tier 1 (Baseline MOQ), Tier 2 (2x to 5x MOQ), Tier 3 (Strategic / High-Volume Reserve). State price validity period.] | [Vendor Guidance: State Incoterm 2020: e.g. FCA (Origin Airport Hub) or CIP Lagos (LOS Airport). State payment milestones: e.g. 30% advance deposit / 70% against shipping documents, or Confirmed L/C.]
Item 1: Medical-Grade Non-Contact Infrared Handheld Thermometers | [State MOQ: e.g. 500 units ¦ Sample availability: 2-5 units on request] | [Tier 1 (500 - 1,999 units): $____ / unit Tier 2 (2,000 - 4,999 units): $____ / unit Tier 3 (5,000+ units): $____ / unit Price Validity: 60 Days] | [Incoterm: FCA (Origin Hub) / CIP Lagos (LOS) Payment: 30% T/T Advance, 70% against AWB & Inspection Certificate / LC at sight]
Item 2: Walk-Through Mass Thermal Imaging Camera System with Blackbody | [State MOQ: e.g. 2 - 5 complete systems] | [Tier 1 (1 - 4 systems): $____ / system Tier 2 (5 - 9 systems): $____ / system Tier 3 (10+ systems): $____ / system Price includes Blackbody calibration & software license] | [Incoterm: FCA / CIP Lagos Payment: Irrevocable Letter of Credit (L/C) at sight or 40% Advance, 60% upon pre-dispatch inspection]
Item 3: Walk-Through Ultraviolet (UV) Disinfection Booth | [State MOQ: e.g. 1 - 2 systems] | [Tier 1 (1 - 2 systems): $____ / system Tier 2 (3 - 5 systems): $____ / system Tier 3 (6+ systems): $____ / system Warranty: 12-month standard included] | [Incoterm: FCA / CIP Lagos Payment: 50% deposit on order, 50% prior to air freight dispatch or Irrevocable L/C]
Item 4: WHO VHF Category III Type 3-B / 4-B Biohazard Coveralls | [State MOQ: e.g. 2,500 units (assorted sizes M, L, XL, XXL)] | [Tier 1 (2,500 - 9,999 units): $____ / suit Tier 2 (10,000 - 24,999 units): $____ / suit Tier 3 (25,000+ units): $____ / suit Price includes sterile packaging] | [Incoterm: FCA (Origin Airport) / CIP Lagos Airport Payment: 30% deposit with PO, 70% balance upon presentation of Bill of Lading / AWB copy]
Item 5: Heavy-Duty Biohazard Mortuary Body Bags (250+ Micron) | [State MOQ: e.g. 1,000 units] | [Tier 1 (1,000 - 2,499 units): $____ / bag Tier 2 (2,500 - 4,999 units): $____ / bag Tier 3 (5,000+ units): $____ / bag Price includes identification tags] | [Incoterm: FCA / CIP Lagos Payment: 30% deposit, 70% against shipping documents / Sight Draft / Irrevocable L/C]
Item 6: Sodium Dichloroisocyanurate (NaDCC) Chlorine Tablets (5g) | [State MOQ: e.g. 50,000 tablets (500 tubs of 100)] | [Tier 1 (50k - 99k tablets): $____ / 100-tab tub Tier 2 (100k - 249k tablets): $____ / 100-tab tub Tier 3 (250k+ tablets): $____ / 100-tab tub Shelf Life Guarantee: 36 Months] | [Incoterm: FCA (Origin Airport) / CIP Lagos Payment: 30% Advance, 70% against Dangerous Goods AWB & Certificate of Analysis]
Item 7: FFP3 / N95 Medical-Grade Biohazard Respirator Masks | [State MOQ: e.g. 10,000 units (25 master cartons)] | [Tier 1 (10,000 - 24,999 units): $____ / unit Tier 2 (25,000 - 49,999 units): $____ / unit Tier 3 (50,000+ units): $____ / unit Packaging: 20 pcs/box] | [Incoterm: FCA / CIP Lagos (LOS) Payment: 30% Advance T/T, 70% against AWB & Certificate of Conformity]

## PQQ5 - Public Sector Track Reco
STRANFIELD TRADING COMPANY ¦ PQQ5: INSTITUTIONAL & PUBLIC SECTOR TRACK RECORD MATRIX
Objective: Assess vendor proven experience in delivering medical and emergency response supplies to public sector bodies, national health ministries, multilateral entities, and international NGOs over the past 3 to 5 years.
Item No. & Equipment Name | Summary of Past Public Sector / Institutional Supply Experience
[Exemplar / Guidance] | [Vendor Guidance: For each item, provide verified details of past contracts: 1) Client / Institutional Agency Name (e.g. Ministry of Health, WHO, UNICEF, Red Cross / IFRC, UK NHS, US CDC); 2) Contract Year & Destination Country; 3) Quantity / Volume Supplied; 4) Operational Context (e.g. Epidemic Outbreak Response, National Strategic Stockpile, Quarantine Deployment); 5) Client Reference / Certificate of Completion available (Yes/No).]
Item 1: Medical-Grade Non-Contact Infrared Handheld Thermometers | [Enter: 1) Client/Agency Name: ____________________ ¦ 2) Year & Country: ________ ¦ 3) Volume: e.g. 15,000 units 4) Deployment: National border screening & hospital outpatient intake ¦ 5) Reference Contact / PoC: Available upon request]
Item 2: Walk-Through Mass Thermal Imaging Camera System with Blackbody | [Enter: 1) Client/Agency Name: ____________________ ¦ 2) Year & Country: ________ ¦ 3) Volume: e.g. 18 systems 4) Deployment: International airport arrivals hall & central government complex ¦ 5) Operational Uptime / Maintenance Support Ref]
Item 3: Walk-Through Ultraviolet (UV) Disinfection Booth | [Enter: 1) Client/Agency Name: ____________________ ¦ 2) Year & Country: ________ ¦ 3) Volume: e.g. 6 installations 4) Deployment: Regional isolation hospital & port quarantine intake facility ¦ 5) Protocol Validation / Commissioning Sign-off]
Item 4: WHO VHF Category III Type 3-B / 4-B Biohazard Coveralls | [Enter: 1) Client/Agency Name: ____________________ (e.g. UN Agency / National Disaster Management) ¦ 2) Year & Country: ________ 3) Volume: e.g. 120,000 suits ¦ 4) Context: Viral epidemic rapid response & treatment center deployment ¦ 5) Inspection / Audit Reference]
Item 5: Heavy-Duty Biohazard Mortuary Body Bags (250+ Micron) | [Enter: 1) Client/Agency Name: ____________________ ¦ 2) Year & Country: ________ ¦ 3) Volume: e.g. 25,000 bags 4) Deployment: National emergency stockpile & hemorrhagic fever containment program ¦ 5) Performance Verification / Proof of Delivery]
Item 6: Sodium Dichloroisocyanurate (NaDCC) Chlorine Tablets (5g) | [Enter: 1) Client/Agency Name: ____________________ (e.g. UNICEF / Red Cross / Ministry of Water & Health) ¦ 2) Year & Country: ________ 3) Volume: e.g. 2,000,000 tablets ¦ 4) Deployment: Emergency cholera / VHF disinfection & municipal field sanitation ¦ 5) Quality Release Report Ref]
Item 7: FFP3 / N95 Medical-Grade Biohazard Respirator Masks | [Enter: 1) Client/Agency Name: ____________________ ¦ 2) Year & Country: ________ ¦ 3) Volume: e.g. 500,000 units 4) Deployment: National strategic healthcare reserve & intensive care clinical stockpile ¦ 5) Government Contract / Tender Award Ref]